## Slide 1
③
①
④
⑧
②
⑤
⑥
⑦
ION
BOX
FIL1_V
FIL2_V
EL3
Q-
Q+
FIL_GND
BODY
⑨
FC/SEM
FIL2
FIL1
ION
BOX
Q_DC+
A
Filament current measure
0.5V/A
meas : max 6A
ion focusing lens
(einzel lens)
+
-
FIL_OS
+
-
AD8236
-
+
REF
IB_OS
1K
1K
Emission
current
LENS_DC
+
-
LENS_DC
+
-
+
-
+
-
RF
Q_DC-
0..+210V
-105..+105V
0..+105V
0..+105V
-105..+105V
FIL_3V
+
-
FIL1_V
FIL2_V
FET
FET
50
0.005
-70V
FIL_OS = IB_OS – 70V
Extracting Lens
Focusing Lens
Accelerating Lens
SEM : -1000V
??? V
(negative ???)
IB_OS
??? V
(positive ???)

## Slide 2
1
2
3
4
5
6
7
8
9
AVC
(RS-485)
1
2
3
4
5
6
7
8
9
Remote IN
(male)
Drive Board(plasma) Block Diagram
MAX14915
24V high side SW
SPI
OUT1
OUT2
OUT3
OUT4
OUT5
OUT6
OUT7
OUT8
Vdd
Va
A0
J15
Flow meter
J16
AUX2
J5
DI
J6
Door SW
J7
AVC
J24
IO
J8
TMP
J21
Plasma PS
J20
DP
J19
CB2.0
J18
Front panel
AD7682(7689)
4ch(8ch) ADC
SPI
IN0
IN1
IN2
IN3
IN4
IN5
IN6
IN7
Vdd
Vio
MAX14915
24V high side SW
SPI
OUT1
OUT2
OUT3
OUT4
OUT5
OUT6
OUT7
OUT8
Vdd
Va
A0
ISO1430
V1
TX
V2
RS485
isolation
RX
A
B
RZ/T1
Renesas MCU
SPI
ENET0
ENET1
ENET2
CLK
MX25L6433
Serial FLASH
SPI
25MHz
MRSSV1DR16SM
rotary selector
1
2
4
8
C
1
2
4
8
C
2.2K
SSR
SSR
P24V
J1 #21
J1 #22
P3V3
P24V
P3V3
MAX31913
HV DI  SPI
SPI
IN0
IN1
IN2
IN3
IN4
IN5
IN6
IN7
Vdd
150
P24V
TP5
TP6
P24V
ECAT IN
P5V
P3V3
P24V
P5V
120
P3V3
ISO1430
V1
TX
V2
RS485
isolation
RX
A
B
120
P24V
TP1
J1 #6
J1 #29
J1 #28
J1 #12
J1 #17
J1 #16
J1#2
J1#3
LAN8710
10/100M PHY
TX RX
RMII
P5V/3A
LMZM33603
(DCDC module)
inductor + FET
LMZ10501
(DCDC module)
inductor + FET
J25
Power
EMI
FILTER
BNX022
0154007.0DR
7A Fuse
24Vdc
R7812-0.5
(DCDC module)
inductor + FET
P12V
P3V3
J1#15
J1#14
1
2
4
8
C
ECAT OUT
ETHERNET
J12
SEM
BYH05-200S01
Vc
Vo+
Vo-
Vin
+200V
P5V
J13
Lens
BYH05-200S01
Vc
Vo+
Vo-
Vin
-200V
J22
Heater-1
SSR
MAX31855
SPI
TC+
TC-
Vdd
P24V
P3V3
24VIN
J14
PKR361
P24V
J26
AUX3
J17
Back panel
P24V
RS-485
J27
DO (Digital Out)
GR
BL
OR
VL
BK
RD
1
2
3
4
5
6
7
8
9
Remote OUT
(female)
GR
BL
OR
YL
BK
RD
2.2K
2.2K
2.2K
2.2K
CB PWR
Plasma PWR
CB to PC conn.
PC to CB conn.
LAN8710
10/100M PHY
TX RX
RMII
LAN8710
10/100M PHY
TX RX
RMII
AC/DC
220Vac
CB power
1
2
3
4
5
6
7
8
9
DP (female)
10
11
12
13
14
15
NTC
GND
24VOUT
Standby
Pump on
DI PRI.
RS-485
(9600, N, 8, 1)
not connected
DP (Diaphragm Pump)
vacuum
exhaust
1
2
3
4
5
6
7
8
9
TMP (female)
10
11
12
13
14
15
GR
BL
YL
BK
RD
WH
SW : TMP test ON
(#2,#5,#7 short)
vacuum
exhaust
TMP
(Turbo Molecular Pump)
J9
FAN1
J10
FAN2
P24V
J11
AUX1
P24V
P24V
Flow meter
FSM-V-AH3-R0005-M5
GR
BL
BK
out : 1~5V
- no flow : 3V
- pos. flow : 3~5V
- neg. flow : 1~3V
J23
Heater-2
SSR
MAX31855
SPI
TC+
TC-
Vdd
P24V
P3V3
HTR1
HTR2
HTR1
HTR2
SEM_ONOFF
Pfeiffer : PKR361
- Pressure gauge
(1x10-8 ~ 1000 hPa)
- out : 2~8.6V log scale
(0.6V/decade)
PKR361 pressure calculation
YL
WH
RD
BK
4
5,6
3
2
BL
GR
2A Fuse
WH
YL
DAC80504
16bit DAC
SPI
OUT0
OUT1
OUT2
OUT3
Vdd
P5V
C11152-01
12V  -1500V
ON/OFF
HV_OUT
HV_MON
Vcc
Vcont
P12V
SEM_ONOFF
-1000V
WH
connection check required
int REF
=4.096V
x0.4
x0.8
x0.8
Chassis
x0.4
SSR : EL100D10-24N

## Slide 3
CB3 PS(power supply) rev 03 Block Diagram
LTC4365
power protection
24VIN
Si7370
N ch MOSFET(60V)
LTM8026
(DCDC module)
inductor + FET
+20V
LTC7149
(DCDC)
+ FET
-20V
LTM8065
(DCDC module)
inductor + FET
+10V
LTM8065
(DCDC module)
inductor + FET
-10V
LTM8065
(DCDC module)
inductor + FET
+7.5V
LTM8065
(DCDC module)
inductor + FET
-7.5V
TPSM53603
(DCDC module)
inductor + FET
+3.6V
IRS-5/10-Q12
(isolated DCDC)
Fan (24V)
FIL_3V, 6A
Filament Power (Isolated)
LT4151
I2C addr : 1101_001
Voltage, Current monitor
1) Sense voltage (for current)
2) VIN voltage
3) ADIN voltage
S+
S-
VIN
ADIN
I2C
R divide
I2C addr : 0101_110
1
2
3
4
5
6
7
8
10
9
11
12
13
14
15
16
17
18
20
19
21
22
23
24
25
26
27
28
30
29
31
32
33
34
35
36
37
38
40
39
41
42
43
44
45
46
47
48
50
49
51
52
53
54
55
56
57
58
60
59
61
62
63
64
65
66
67
68
70
69
71
72
73
74
75
76
77
78
80
79
5A
4A
2.5A
2.5A
2.5A
2.5A
3A
+7.5V
+10V
+24V
GND
+3.6V
J4
(FAN connector)
GND
+20V
-20V
-10V
-7.5V
FIL_3V
FIL_GND
J2 : to IDBB
+3.3VIN
MIC2774L-46YM5
power supervisor
1) +5V > 4.68V
2) +20V > 17V
IN
Vcc
MIC2774
RST
MR
+5VIN
GND
+20V
464K
8.25K
+5VIN
(32Kbyte)
Vcc
I2C
AT24C256
I2C addr : 1010_000
+3.3VIN
use PGOOD signal + pull-up
1 : Power Good
0 : Power fail
I2C_SCL
I2C_SDA
damping : 100 ohm
pull-up : 10K
I2C
PWM
TACH
Vcc
EMC2103
(FAN controller)
ALERT
I2C_SCL
I2C_SDA
LED
ISO_5V
10A
LT8612
(DCDC)
+ FET
LED
PG
U3
(N/A)
FAN_ALERT-
+24V
+24V
FIL_GND

## Slide 4
ATIK IDBB rev0.1 Block Diagram (1/3)
1
2
3
4
5
6
7
8
10
9
11
12
13
14
15
16
17
18
20
19
21
22
23
24
25
26
27
28
30
29
31
32
33
34
35
36
37
38
40
39
41
42
43
44
45
46
47
48
50
49
51
52
53
54
55
56
57
58
60
59
61
62
63
64
65
66
67
68
70
69
71
72
73
74
75
76
77
78
80
79
+7.5V
+10V
+24V
GND
+3.6V
+20V
-20V
-10V
-7.5V
FIL_3V
J18 : to PS
+3.3VIN
I2C1_SCL
I2C1_SDA
LED
U3
(N/A)
+5VIN
FIL_GND
= FIL_OS
PGOOD
FAN_ALERT-
1
2
3
4
5
6
7
8
10
9
11
12
13
14
15
16
17
18
20
19
21
22
23
24
25
26
27
28
30
29
51
52
53
54
55
56
57
58
60
59
61
62
63
64
65
66
67
68
70
69
71
72
73
74
75
76
77
78
80
79
GND
J33 : to RFB
41
42
43
44
45
46
47
48
50
49
91
92
93
94
95
96
97
98
100
99
1
2
3
4
5
6
7
8
10
9
11
12
13
14
15
16
17
18
20
19
21
22
23
24
25
26
27
28
30
29
31
32
33
34
35
36
37
38
40
39
41
42
43
44
45
46
47
48
50
49
51
52
53
54
55
56
57
58
60
59
61
62
63
64
65
66
67
68
70
69
71
72
73
74
75
76
77
78
80
79
GND
???
J9 : to FP
31
32
33
34
35
36
37
38
40
39
81
82
83
84
85
86
87
88
90
89
+3.6V
-20V
-7.5V
+20V
+7.5V
+3.3VIN
GND
+5VIN
GND
-10V
+3.6V
+10V
DAC1_SYNC
DAC1_SCLK
DAC1_SDI
DAC1_SDO
I2C1_SCL
I2C1_SDA
FAN_ALERT-
J18_42
J18_43
J18_44
J18_45
J18_42
LENS_DC
IONBOX
I2C2_SCL
I2C2_SDA
FP_08
FIL2_V
FIL_GND
FIL1_V
WP
FP_01
FP_02
FP_07
FP_04
FP_05
RESET-
FP_06
FP_03
T_alert-
FP_15
FP_14
FP_13
FP_12
FP_11
FP_10
FP_09
SCL_1
SDA_1

## Slide 5
ATIK IDBB rev0.1 Block Diagram (2/3)
LT3062
(LDO)
9V
LT3062
(LDO)
5VA
LM337
(LDO)
-5V
+10V
+7.5V
-10V
LT3062
(LDO)
F_P5V
F_P9V
F_N9V
(unused)
VB_P9V
VB_N9V
LT3062
(LDO)
VB_P5V
LT3062
(LDO)
VB_P1.4V
VB_N3.6V
+24V
+110V
-110V
+215V
MC33161
Vcc
IN1
IN2
OUT1
OUT2
MC33161
Vcc
IN1
IN2
OUT1
OUT2
under voltage
monitor
+3.3V
5V
+3.3V
(105V)
(207V)
(-105V)
HV_FAIL
+3.3V
5VA
0V
AD5676
Vdd
SPI
Vlogic
Vout0
Vout1
Vout2
Vout3
Vout4
Vout5
Vout6
Vout7
Vref
GAIN
(8ch, 16-bit DAC)
RESET
J33
+215V
5V
PCA9306
V1
I2C1
V2
I2C2
MCP3424(4ch ADC)
Vdd
I2C
CH1
CH2
CH3
CH4
MCP3424(4ch ADC)
Vdd
I2C
CH1
CH2
CH3
CH4
MCP3424(4ch ADC)
Vdd
I2C
CH1
CH2
CH3
CH4
ADHV4702
+
-
+110V
-110V
2K
100K
Vo=51(2.048-V1)
ADHV4702
+
-
+110V
-110V
2K
100K
Vo=51(2.048-V1)
QDCC+
IONBOX_DCC
LENS_DCC
FILN_REF
FIL_OFFSETC
QDCC-
QDCN_REF
DAC_VREF_ADJ
Q_DC+
OP2177
-
+
VREF_2048 or QDCN_REF
5V
0V
ADHV4702
+
-
+215V
-5V
2K
100K
Vo=51V1
Q_DC-
OP2177
-
+
VREF_2048or FILN_REF
5V
0V
FIL_OS
ADHV4702
+
-
+110V
-5V
4K
100K
Vo=26V1
ADHV4702
+
-
+110V
-5V
4K
100K
Vo=26V1
IONBOX_OS
LENS_DC
OP2177
-
+
5V
0V
OP2177
-
+
VREF_4096
5V
0V
differential amp
DAC1_VREF (4.096V)
VREF_2048
: R divide
Q_DC+
VREF(4.096V)
+110V
IONBOX_OS
LENS_DC
VREF(4.096V)
VREF(4.096V)
-110V
ratio=0.008, 0~+265V
-2.048V~+2.048V, 18-bit ADC
ratio=0.332, 0~+6V
ratio=0.008, 0~+256V
unused
Q_DC-
ratio=0.0175, -115~+115V
ratio=0.008, 0~+256V
FIL_OS
ratio=0.0175, -115~+115V
ratio=0.332, 0~+6V
ratio=0.0175, 0~+115V
ratio=0.0175, 0~+115V
ratio=0.25, 0~+8V
ratio=0.0175, -115~+115V
+3.3V
MEV1D2409SC
(isolated DCDC)
MEV1D2409SC
(isolated DCDC)
ISC1540
V1
I2C1
V2
I2C2
isolation
U136
ISC1540
V1
I2C1
V2
I2C2
isolation
U16
+3.3V
I2C3
I2C2
I2C_FDAC
F_P5V
MCP4728
(4ch 12-bit DAC)
Vdd
I2C
VoutA
VoutB
VoutC
VoutD
BUSY
DAC1_VREF
I2C2
FIL_3V
AD8572
+
-
FIL_3V
Filament
1 (#8)
x1
x100
MCP3424(4ch ADC)
Vdd
I2C
CH1
CH2
CH3
CH4
MCP3424(4ch ADC)
Vdd
I2C
CH1
CH2
CH3
CH4
FIL1_V
AD8572
+
-
Filament
2 (#3)
x1
x100
FIL2_V
ratio=0.5
ratio=0.5
FIL_3V
ratio=0.5
FREF_4096
ratio=0.33
F_P5V
MAX6341
4.096Vref
FREF_4096
FREF_2048
½
½
2.5V
F_CURRENT
ratio=1.0
unused
FREF_4096
ratio=0.33
F_P5V
AD8236
-
+
REF
5m
FREF_2048
F_CURRENT
(1A  0.5V)
G=100
I2C:1101_100
I2C:1101_011
I2C:1100_000
Possible to use one I2C isolator
I2C:1101_010
MAX6341
4.096Vref
VBREF_0496
LM337
(LDO)
(32Kbyte)
Vcc
I2C
AT24C256
I2C:1010_xxx
ISC1540
V1
I2C1
V2
I2C2
isolation
U18
I2C1
VB_P5V
+3.3V
MCP3424(ADC)
Vdd
I2C
CH1
CH2
CH3
CH4
(18 bit, 2.048Vref)
ADS1115(ADC)
Vdd
I2C
CH1
CH2
CH3
CH4
RDY
(16 bit ΔΣ)
I2C:1101_101
I2C:1001_011
ISC1540(U137)
V1
I2C1
V2
I2C2
isolation
ISC1540(U101)
V1
I2C1
V2
I2C2
isolation
EMISSION_CUR
EMISSION_CUR
VB_P5V
VBREF_0496
VB_P5V, N3.6V
VB_N3.6V
PS connector
+20V
-20V
+10V
-10V
+7.5V
-7.5V
+3.6V
GND
+24V
FIL_GND
FIL_3V
+3.3V
+5V
I2C_PS
PGOOD/FAN_ALERT-
LED
LTM8024
(DCDC module)
+FET +inductor
+5V
+3.3V
LED
+24V
+24V
LT8365
(DCDC)
+FET
EN
LT8365
(DCDC)
+FET
EN
LT8365
(DCDC)
+FET
EN
MAX6341
4.096Vref
(0~+208.9V)
(-100.35V~
+104.45V)
(-100.35V~
+104.45V)
(0~+106.5V)
(0~+106.5V)
DAC1_SPI
VB_P1.4V
AD8236
-
+
REF
what’s the purpose of this block ?
same measurement with MCP3424
G=5
VB_N3.6V
VB_GND
VB_GND
50
IONBOX
J33 (side 51)
+5V
+3.3V
from FP
FIL_3V
FIL_GND
FIL_GND
FIL_GND
int. ref=2.048V
Gain=x1/x2
FIL_GND
unused
Filament Block
I2C_FADC
FIL1_V
FIL2_V
VB_GND
ratio=0.5
ratio=0.33
ratio=0.33
EMISSION_CUR
ratio=0.5
VBREF_0496
ratio=0.33
ratio=0.5
IONBOX Block
out range : 0 ~ 4.096V
VREF_4096
I2C:1101_001
I2C:1101_000
ADC required items
1) Q_DC+
2) Q_DC-
3) +215V
4) +110V
5) -110V
6) FIL_OS
7) IONBOX_OS
8) LENS_DC
ADC required items
1) FIL1_V
2) FIL2_V
3) F_CURRENT
4) FIL_3V
1/2

## Slide 6
ATIK IDBB rev0.1 Block Diagram (3/3)
+215V
-5V
+
-
ADHV4702
12K
12K
2K
100K
1K
Vo
Vb
Va
Non-inverting Amp
Va=Vb=V1, i1 = Va/R1
Vo = i1*R2 + Va
= Va*R2/R1 + Va
= (1+R2/R1) V1
= 51 V1
i1
R1
< Q_DC+ circuit >
Non-inverting amp configuration
1) input : 16-bit DAC voltage
0 ~ +4.096V, 1 step = 62.5uV
2) output : Vo=(1+R2/R1)Vi = 51Vi
0 ~ +209V, 1 step = 3.1875mV
DAC OUT5
(0~4.096V)
+110V
-110V
+
-
ADHV4702
12K
12K
2K
100K
1K
Vo
Vb
Va
Differential Amp
Va=Vb=V2=2.048V
i1 = (Va-V1)/R1
Vo = i1*R2 + (Va-V1)
= (Va-V1)*R2/R1 + (Va-V1)
= (1+R2/R1) (Va-V1)
= 51 (2.048 – V1)
i1
R1
< Q_DC- circuit >
Differential amp configuration
1) input : V+ : fixed 2.048V
V- : 16-bit DAC voltage
0 ~ +4.096V, 1 step = 62.5uV
2) output : Vo=(1+R2/R1)(V2-V1) = 51(V2-V1)
-104.45 ~ +104.45V, 1 step = 3.1875mV
VREF / 2
= 2.048V
V1
V1
V2
R2
R2
DAC OUT0
(0~4.096V)
Q_DC+
(0~+209V)
Q_DC-
(-104.45 ~
+104.45V)
Vi
Vo
0
+210
1
2
3
4
Vi
Vo
0
+110
-110
1
2
3
4
DAC OUT4
(0~4.096V)
+110V
-110V
+
-
ADHV4702
12K
12K
2K
100K
1K
Vo
Vb
Va
Differential Amp
Va=Vb=V2=2.048V
i1 = (Va-V1)/R1
Vo = i1*R2 + (Va-V1)
= (Va-V1)*R2/R1 + (Va-V1)
= (1+R2/R1) (Va-V1)
= 51 (2.048 – V1)
i1
R1
< FIL_OS circuit >
Differential amp configuration
1) input : V+ : fixed 2.048V
V- : 16-bit DAC voltage
0 ~ +4.096V, 1 step = 62.5uV
2) output : Vo=(1+R2/R1)(V2-V1) = 51(V2-V1)
-104.45 ~ +104.45V, 1 step = 3.1875mV
VREF / 2
= 2.048V
V1
V2
R2
FIL_OS
(-104.45 ~
+104.45V)
Vi
Vo
0
+110
-110
1
2
3
4
+110V
-5V
+
-
ADHV4702
12K
12K
4K
100K
1K
Vo
Vb
Va
Non-inverting Amp
Va=Vb=V1, i1 = Va/R1
Vo = i1*R2 + Va
= Va*R2/R1 + Va
= (1+R2/R1) V1
= 26 V1
i1
R1
< IONBOX_DC, LENS_DC circuit >
Non-inverting amp configuration
1) input : 16-bit DAC voltage
0 ~ +4.096V, 1 step = 62.5uV
2) output : Vo=(1+R2/R1)Vi = Vi
0 ~ +106.5V, 1 step = 1.625mV
V1
R2
DAC OUT2
DAC OUT3
(0~4.096V)
IONBOX_DC
LENS_DC
(0~+106.5V)
Vi
Vo
0
+110
1
2
3
4

## Slide 7
CB3 RFD rev0.2 Block Diagram (1/2)
+7.5V
MAX16025
Vcc
IN1
IN2
OUT1
OUT2
MR-
RESET-
FPGA
Artix 7
XC7A35T
MAX5885
16-bit DAC, 200M
FS current : 10mA
@50Ω : 0.5Vpp
IOUT+
IOUT-
D[15:0]
AVdd
DVdd
CLK
TPS7A4701
(LDO)
+7V, 1A
LT3029 U15
OUT1
OUT2
-7.5V
TPS7A33
(LDO)
-7V, 1A
+20V
-20V
+3.6V
LT3029 U18
OUT1
OUT2
3.3V1, 500mA
3.3V2, 500mA
1.8V1, 500mA
3.3V3, 500mA
LT3029 U21
OUT1
OUT2
1.8V2, 500mA
3.3V4, 500mA
LT3045
(LDO)
1.0V, 500mA
+1.0V, +1.8V1, 3.3V3
3.3V2
power mon
50MHz ↑
50M : 14.3 step
70M : 20 step
100M : 28.6 step
140M : 40 step
200M : 57 step
DDS
LT3032-5
(dual LDO)
+5V, 150mA
-5V, 150mA
+5V
+7V
-7V
+
-
-
+
THS6212
Line Driver
G = 10
17T
70T
N=4.12
③
①
④
⑧
②
⑤
⑥
⑦
EL0
FIL_1
FIL_2
EL3
Q_pole-
Q_pole+
FIL_GND
+5V
-5V
+
-
LT6411
ADC driver
G = 2
+
-
IN+
IN-
D[15:0], OF
OVdd
Vdd
CLK
LTC2165
16-bit ADC
(100 M)
+1.8V2
SENSE
TP31,32
100M_ADC
SENSE+
Vin
SENSE-
ADIN
20m
I2C
LC
Filter
serial FLASH
Vcc
DQ
MT25QL128
EEPROM 32K
I2C: 1010_001
Vcc
I2C
AT24C256
WP
LTC6655
Vcc
1.25V
1.25V ref
+5V
SPI
temp. sensor
I2C: 1001_010
Vcc
I2C
TCN75A
Alert
RF=3.5MHz
J1
J5
J2
J3
LTC4151
I2C addr : 1101_001
Voltage, Current monitor
1) Sense voltage (for current)
2) VIN voltage
3) ADIN voltage
3.3V3
3.3V1
CLK1
CLK2
clock generator
CLK3
CLK4
CLK
Vdd
I2C
5P49V5901
J10 conn.
X-tal
(25MHz)
3.3V4
where ???
0~0.5V
@50Ω
Vcc/2
biased
PRE
AMP
OUT
AMP
AD8370 (VGA)
SER
Vcc/2
biased
DATA, CLK, LATCH
0~+1V
0~+4V
J4
0~+18V
0~-18V
ADA4870
+
-
+20V
-20V
300
1.2K
G = 5
ON
SD
TFL
0~-4V
ADA4870
+
-
+20V
-20V
300
1.2K
G = 5
ON
SD
TFL
-20V
0~-1V
AC divider
-1 ~ +1V range
For 2V input range
VCM(0.9V)±0.5V
= 0.4 ~ 1.4V
ADC control signal
IONBOX
LENS_DC
FIL2_V
FIL1_V
NA
Q_DC+
Q_DC-
1
2
3
4
5
6
7
8
10
9
11
12
13
14
15
16
17
18
20
19
21
22
23
24
25
26
27
28
30
29
51
52
53
54
55
56
57
58
60
59
61
62
63
64
65
66
67
68
70
69
71
72
73
74
75
76
77
78
80
79
GND
J33 : to IDBB
41
42
43
44
45
46
47
48
50
49
91
92
93
94
95
96
97
98
100
99
31
32
33
34
35
36
37
38
40
39
81
82
83
84
85
86
87
88
90
89
+3.6V
-20V
-7.5V
FIL2_V
FIL_GND
+20V
+7.5V
FIL1_V
SCL_1
SDA_1
I2C_1
I2C1
(3.3V1 pull-up)
T_alert-
T_alert-
WP
WP
FP_01
FP_02
FP_07
FP_04
FP_05
RESET-
RESET-
FP_15
FP_14
FP_13
FP_12
FP_11
FP_10
FP_09
FP_08
FP_06
FP_03
3.3V3
NA
FP_01 ~ FP_15
I2C:1101_010
max 1Vpp differential
max gain = x7(LG), x50(HG)
used gain range = 0~x4
0~+0.25V
0~-0.25V
VC
40pF
Quadrupole
capacitance
LC resonant frequency =

## Slide 8
17T
70T
N=4.12
RF transformer
primary : 17T, AWG27(Ø0.36)
secondary : 70T, AWG27(Ø0.36)
N = 70T / 17T = 4.12
Lodestone Pacific
- HTM851-6 (Digikey)
(Horizontal Toroid Mount)
- 4 pin 짜리는 찾지 못 함
Toroidal Ferrite Core
- Micrometals : T80-2 (레오콤)
- size : 20.2 x 12.6 x 6.35
AWG28 enamel coated wire
- 28H200P.125 (Digikey)
(AWG28 enamel coated wire)
Toroidal split holder
- Can’t find similar model
- make RF transformer
without this holder
Toroidal core inductance calculation
Turns =  L(nH) =
< 계산값 > < LCR meter 측정값 > < 실험값 >
1) primary : L = 5.5 = 1.6 uH L =5 uH L =2.5 uH
Z = 2πfL = 35 Ω Z = 110 Ω Z = 55 Ω
2) secondary : L = 5.5 = 27 uH L = 50 uH L = 23.4 uH
Z = 2πfL = 593 Ω Z = 1100 Ω Z = 514 Ω
Note :
1) Toroidal ferrite core 는 size 와 color 를 기준으로 검색 (다를 수 있음).
2) Toroidal core 의 bobbin(case, holder) 은 유사 제품 찾을 수 없음.
자체 제작시 bobbin 없이 절연 taping 후 wire 를 감아서 제작.
3) transformer 의 1차, 2차 측 inductance 계산 값이 LCR meter 측정값과 다름.
- 계산값, LCR meter측정값, 실험값은 아래 내용 참조
1) 계산값은 ideal 한 상태의 값이므로 실제 측정치와 차이가 발생함
2) LCR meter 측정값은 freq. = 100KHz 에서의 측정치
3) 실험값은 실 사용 조건과 가장 유사한 상태에서 측정.
(freq. = 3.5 MHz, signal = 1Vpp sinewave)
3M insulation tape 로 대치
- 3M159698-ND (Digikey)
- width : 6.35mm
- core 에 절반씩 겹치도록 taping
17T
70T
+
-
1Vpp @3.5MHz
(sinewave)
V1
V2
55Ω
1차측 inductance 실험
V1 = V2
+
-
1Vpp @3.5MHz
(sinewave)
V1
V2
514Ω
2차측 inductance 실험
70T
17T
V1 = V2
CB3 RFD rev0.2 Block Diagram (2/2)

## Slide 9
CB2.0 Ion Detector Board : rev A
LT3045
U19
(LDO)
5V
6.7V
LT3045
U51
(LDO)
5.6V
LT3045
U50
(LDO)
3.3V2
LT3045
U49
(LDO)
2.5V
ADA4530
FC / SEM
LTC2380-24
24bit ADC
IN+
IN-
REF
Vdd
OVdd
SDI
SPI
SCK
SDO
BUSY
CNV
2.5V
+
-
AD8066
+
-
Si8652
(isolator)
Vdd2
Vdd1
isolation
ISO7341
(isolator)
Vdd2
Vdd1
isolation
24AA512
(64Kbyte)
Vcc
I2C
WP
TCN75A
(temp. sensor)
Vcc
I2C
to FP board
ADG1613
5V
3.3V1
3.3V1
power (???)
3.3V1
3.3V1
3.3V2
3.3V2
J1
GND1
3.3V1
6.7V
I2C_SDA
I2C_SCL
EE_WP
TP139
TP136
TP137
TP138
ADC_SDI
ADC_SDO
ADC_BUSY
1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
ADC_CNV
ADC_SCK
ADC_SDI
ADC_SDO
ADC_BUSY
3.3V1
10K
I2C_AD : 100_1000
I2C_AD : 101_0000
LTC6655
(5V ref.)
Ref5V
Ref5V
+
-
+
-
GND2
10K
10K
Ref2.5V
unused
AD8066
+
-
R206
51K
C72
390p
C283
680p
5V
Ref2.5V
GND2
C71
100p
R203
51K
R204
10K
R205
10K
R20
499
R21
499
GND2
GND2
C271
10n
C272
100n
C270
10n
GND2
C282
5p
R958
1M
R967
100M
R959
0
R59, 33
R58, 33
GND2
R28, 33
GND1
GND1
GND2
R11, 0
GND1
U21
LC filter
FC : Faraday Cup
SEM : Secondary Electron Multiplier
ADC_CNV
ADC_SCK
100p
Tow-Thomas Biquad Filter (Low Pass Filter)
- Fc = 4.5 KHz
Low Pass Filter
- Fdiff = 1 / 2πR2Cdiff = 1.6 KHz
- Fcmn = 1 / 2πRCcmn = 32 KHz
Murata : NFE61HT102
- LC Filter
390p//680p
= 250p
two pole compensation
Notes: Femto : 10^(-15) ,, Mili  Micro  Nano  Pico  Femto
250fA  250uA ( 10^(9)) 10억배

## Slide 10
Tow-Thomas Biquad Filter
- reference : History and progress of the Tow–Thomas bi-quadratic filter. by Ahmed M. Soliman
Derivation of loop gain and stability test for low-pass Tow-Thomas biquad filter by Minh Tri Tran
- derivation of 2nd order active biquad filter, 2 filter outputs (LowPass, BandPass)
sL
1. Basic Form Generation : using passive RLC low pass filter
2. Add Parameters : basic form + gain/Q control parameter
- a : Q factor control
- b : gain control
e1=
e2=e1
Fc =
(2) Q =
(3)
250p
100p
51K
Fc adj.
10K
1M
10K
(Q adj.)
-1
VI
VLP
+
+
lossy integrator
lossless integrator
e1
e2
-1
VI
VLP
+
+
b
gain control
Q control
VI
VLP
+
-
C1=250p, C2=100p, R1=80K, R2=51K, R3=1M
1) Fc =
2) Q = 56
3)
C1=100p, C2=100p, R1=76K, R2=51K, R3=1M
1) Fc =
2) Q = 34
3)
Notes: Femto : 10^(-15) ,, Mili  Micro  Nano  Pico  Femto
250fA  250uA ( 10^(9)) 10억배

## Slide 11
OP-Amp frequency compensation : 2-pole compensation
Z
R2
C1
C2
Vi
390p
680p
51K
impedance calculation
Z1 = XC1 // R2 =
Z = Z1+XC2 = =
divide SC2 =
ADA4530
5V
+
-
R2
C1
C2
Vi
V1
R1
Vo
for inverting amp,
G=K, H= , H-1=
= = -K
- reference : Dynamic Circuit Response. by Dennis L. Feucht
- OP-amp compensation required to prevent oscillation at high frequency
1) dominant pole compensation : commonly used, loop gain rolls off(-20dB/dec) at low freq. and reduces loop gain at higher freq.
2) 2-pole compensation : rarely used, loop gain rolls off (-40dB/dec) at higher freq. and then revert to -20dB/dec at Fzero freq.
dominant pole compensation
C
R
impedance calculation
Z = XC // R =
Fz Fu preferable
Fz Fu / 50 = 2MHz / 50 = 40 KHz
two pole compensation
Gain
K
Fd
Fc
Fz
Fu
dominant pole
compensation
(-20dB/dec)
two pole
compensation
(-40dB/dec)
For Tow-Thomas filter R1 calculation
use following assumption
1) TT_R1 = R2(1+C1/C2)
= 51K (1+390p/680p) = 80 K
2) TT_R1 = R2(1+C1/C2)
= 51K (1+150p/300p) = 76 K
C1=390p, C2=680p, R2=51K, C1//
1) Fz =
= = 2.9 KHz
2) Fc = ,  gain(K) ↑ : Fc ↓
= 1.4 KHz
C1=150p, C2=300p, R2=51K, C1//C2=100p
1) Fz =
= = 7 KHz
2) Fc = ,  gain(K) ↑ : Fc ↓
= 3.3 KHz
Notes: Femto : 10^(-15) ,, Mili  Micro  Nano  Pico  Femto
250fA  250uA ( 10^(9)) 10억배